## Sheet1
NOMAC-Operations Excellence Platform
no# | Plant | Observation | Suggestion/Recommendation | Priority | Target Date | Remarks/Response from the site | Status | Reference
1 | Al-Dur2 | Increased ambient temperatures in the back part of the container (hot section). The warm sample blowdown evacuation line was warm, which can indicate a passing valve | • Identify passing valves on sample blowdown lines. | High | May 2, 2023 |  | Open
2 | Al-Dur2 | The thermal shutoff valve and pressure regulator are installed after the secondary cooler | Assess the safety aspect of the TSV location | High | May 2, 2023 |  | Open
3 | Al-Dur2 | The container entrance is unsafe; the riser is too high, and the concrete slab is too narrow.No emergency door in the container | Assess safety issues with the HSE. | High | May 2, 2023 |  | Open
4 | Al-Dur2 | Filters installed on the sample line protect the instruments from possi-ble crud (oxide particles). However, the placement of the filters never allows for measuring the total iron concentration. Iron oxide particles retained are not measured. | Add to the chemistry manual that larger iron oxide particles might not reach the grab sample point | High | May 2, 2023 |  | Open
5 | Al-Dur2 | After some measurements, the free outflow of the sample is blocked by a drain line back-pressure issue. Sodium measurements are, therefore, unavailable. | Take immediate measures to ensure the sample outflow on the sodium analyzers, even if temporary. Siemens does not allow the startup and operation of the steam turbine without sodium measurement. Further-more, condensate sodium measurements are a must for plants with seawater cooling. | High | May 2, 2023 |  | Open
6 | Al-Dur2 | pH measurements without Process Temperature Compensation (PTC) activated do not give the correct result when the sample temperature deviates from 25 °C. | Verify if the pH PTC is set correct on every transmitter. If not, modify and program -0.030 pH/°C as PTC. | High | May 2, 2023 |  | Open
7 | Al-Dur2 | some HK-338 conductivity analyzers, the temperature compensation is not adapted to the nature of the sample. | Correct the temperature compensation on the HK-338 conductivity transmitters: 2% for SC and 1.4% for CC. | High | May 2, 2023 |  | Open
8 | Al-Dur2 | The Polymetron 9523, used to measure the SC, CC, and Degassed Con-ductivity (DC), had the temperature compensation set to linear 2% for all 3 parameters although the parameters are different. | Correct the temperature compensation on the Polymetron 9523 instru-ments (see picture 4): Select "ammonia" for SC, and "HCl" for CC and DC. | High | May 2, 2023 |  | Open
9 | Al-Dur2 | Sodium analyzers, key instruments, were not in service due to sample flow issues | Bring back into the service | High | May 2, 2023 |  | Open
10 | Al-Dur2 | Key Sodium Online Analyzer is not available at Condensate Exctarction Pump outlet | Sodium Analyzers to be installed | Moderate | Dec 1, 2019 |  | Closed | REP-2019-NRR002
11 | Hassyan | Key Online Instruments are missed | Cation Conductivity online analyzers are requried at Reheat steam & HP heater drain | Moderate | Jan 31, 2020 | As suggested, EPC agreed to provide the instruments | Closed | REP-2019-NHY001
12 | Hassyan | Super heated Steam Sodium analyzer | Sodium value instrument has to be verified with PC | High | Aug 30, 2020 |  | Open | HL- MEM-2020-NHY105
13 | Ibri | Ammonia dosing too low for plants with ACC | Increase ammonia dosing to 17 µS/cm measured on sample downstream ammonia dosing point | High | Nov 1, 2019 | Amonia dosing increased to SC 11.2 and pH 9.7 (Condensate) | Closed | REP-2019-NIB001
14 | Ibri | Ammonia dosing is best on auto control given the shifting load | Ammonia dosing pump fixed stroke - Frequency controlled | High | Sep 30, 2020 | Discussoin in progress with EPC to Commission the system | Open | REP-2019-NIB001
15 | Ibri | Dosing TSP | No TSP in the HP drums. | High | Nov 1, 2019 | implemented | Closed | REP-2019-NIB001
16 | Ibri | Process Temperature compensation must be activated on all pH analyzers | Modify the instrument settings | Moderate | Nov 15, 2019 | In progress | Closed | REP-2019-NIB001
17 | Ibri | Some instrument s analog output signals need modifications | Modify instrument analog output signal according report REP-2019-XXX002 | Moderate | Dec 31, 2019 | In progress | Closed | REP-2019-NIB001
18 | Ibri | Reference Instruments | Purchase reference instruments for conductivity and pH | Low | May 1, 2020 | Due to covid-19 pandemic delivery is due | Closed | REP-2019-NIB001
19 | Ibri | Upgrade the lab work | Use of reference instruments for conductivity and pH instead of grab sample analysis | Low | Jun 30, 2020 | To be completed when points no 6 is completed | Closed | REP-2019-NIB001
20 | Ibri | Instrument validation program | Implement the Instrument Standard Operating Procedures | Low | Jun 30, 2020 | To be completed when points no 6 is completed | Closed | REP-2019-NIB001
21 | Ibri | No written procedures available addressing upset conditions | Write action program for high/low pH condensate, high/low pH LP/IP drum water, high silica, high CC in HP drum water | Moderate | May 1, 2020 | In progress | Closed | REP-2019-NIB001
22 | Ibri | No written procedures available addressing instrument nonconformity | Write procedure objectively outlining when an instrument needs maintenance or calibration by I&C department | Moderate | Jun 30, 2020 | In progress | Open | REP-2019-NIB001
23 | Noor 1 | Damage of Sodium On line analyzer | Change the damage parts by a new spare or Buy a new Sodim Analyzer (for condenser) | High | Jan 1, 2020 | Completed | Closed | REP-2019-NRR002
24 | Noor 1 | Damage of Silica On line analyzer | Change the damage part by a new ones (or buy a new silica analyzer for Steam) | High | May 31, 2020 | Completed | Closed | REP-2019-NRR002
25 | Noor 1 | SWAS circuit damaged (valve leaks, damaged sensors, colomn resin, flowregulation...) | Change damage pipeline, resin column & sensor | Moderate | Jun 30, 2020 | Completed | Closed
26 | Noor 1 | The Dual pH analyzers on the sampling system are not suitable to the nature of the samples and the process (SG drum, ST Steam, Condenser) | Replace The dual instruments by the correct single model allowing for process temperature compensation | Moderate | Jul 30, 2020 | Pending | Open | REP-2019-NRR002
27 | Noor 1 | Lack of proper validation of WSC Online analyzers ( ph meter; cond; oxygen) | Implement a modern instrument standard operating procedure using portable reference instruments for the validation of the online monitoring equipment. | Moderate | Apr 30, 2020 | Completed | Closed | REP-2019-NRR002
28 | Noor 1 | Lose control of sample sequencer | Repair or buy new touch screen of sample sequencer buy Sequencer of the same brand of online analyzer | Moderate | Feb 29, 2020 | Pending | Open
29 | Noor 1 | High Value of CO2 in Demi water Tank & Condesate make up tank | Replace the CO2 Absorber with new one Install new CO2 Absorber system | Moderate | Aug 29, 2020 | Pending | Open
30 | Noor 1 | High Value of CC in condeser sample during night | Install a new CO2 degasser to mesure the DCC in condenser sample | Low | May 15, 2020 | Completed | Closed
31 | Noor 1 | The FW pH analyzer gave incorrect too high results and Dosing of Ammonia basing in the Feedwater pH | Configurate dosing system basing in Feedwater computed pH (by using SC & CC) | Low | Jul 15, 2020 | Pending | Open | REP-2019-NRR002
32 | Noor 1 | Unavailability of dosing pump in case of failure | Purchase spare part or complete dosing pump | Moderate | Aug 30, 2020 | In progress | Open
33 | Noor 1 | WSC chemistry monitoring procedure not cover action related to Modifying the dosing of the treatment products | Review WI NNR1-OPA WI-105 and add the dosing philosophy into SGS | Low | Dec 1, 2019 | Completed | Closed | REP-2019-NRR002
34 | Noor 1 | WSC chemistry monitoring procedure not covering the actions related to Start of the steam turbine & action related to operation stuff | WI in place describe actions to be taken in case of chemistry deviation from the normal level (refer to NNR1-OPA WI-105). The deviation must be related to results from the on-line instruments. | Low | Dec 1, 2019 | Completed | Closed | REP-2019-NRR002
35 | Noor 1 | WSC chemistry monitoring procedure not covering actions related to Emergency stop of the plant in case of condenser leak. | WI in place for Condenser Leak ( NNR1-OPA WI-106) | Low | Dec 1, 2019 | Completed | Closed
36 | Noor 1 | WSC chemistry monitoring procedure not covering actions related to Preservation of the cycle during idle periods. | WI in place for SGS Preservation (NNR1-OPA WI-107) | Low | Dec 1, 2019 | Completed | Closed
37 | Noor 1 | The WSC chemistry guidelines of ST manufacturing is general | Provide a guideline adequate to Solar plant (indicate specific value for the startup, variation of load during day) | Low | May 15, 2020 | In progress | Open
38 | Noor 1 | Unavailability of sodium hypochlorite injection in raw water tank and service water tank | repair the leaks in injection pipeline and repair th dosing pumps | Low | Apr 30, 2020 | In progress | Open
39 | Noor 1 | pretreatment is out service no coagulant and flocculant add to raw waterprior to DAF and multimedia filters | perform jar test and put in service coagulation and floculation dsoinf system | Low | Oct 30, 2020 | In progress | Open
40 | Noor 2 | Microbilogical specifications of Raw water not include in WTP manuals | APO should provide Microbilogical specifications limits of Raw water | Moderate | Jul 30, 2020 | according to APO no Raw water quality specifications in RFQ of WTP between RW supplier (MASEN) and owner (APO) | Closed
41 | Noor 2 | unavailability of Deahydrator in Auto mode | configurate Auto mode for Dehydrator | Low | Jul 30, 2020 | send as defect and according to EPC the auto mode not including in RFQ ,and with rarely using of system manual operation | Closed | 06GBS20AT002
42 | Noor 2 | ACF 22 inlet folw meter display wrong value | repair /change the flow meter | Moderate | Jul 30, 2020 | PO for procurement of new flow meters had been send by I&C team | Closed | 06GCB22CF001
43 | Noor 2 | Historic Archiving and Database Backup not include All parametres | Database Backup not include sould include all parametres | High | Jul 30, 2020 | include in WTP improvemenet plan and implementaion ongoing by ROCKWELL AUTOMATION | Closed
44 | Noor 2 | Lack of trending interface in WTP DCS | implementation of trend interface with possibility to display many parametres in one interface | Low | Jul 30, 2020 | include in WTP improvemenet plan and implementaion ongoing by ROCKWELL AUTOMATION | Closed | REP-2019-NRR002
45 | Noor 2 | DM silica analyzer not operational require change spart part | Repair /calibration of DM silica analyzer | High | Jul 30, 2020 | ongoing with EPC to change preheated card and comissioning of appartus | Open | REP-2019-NRR002
46 | Noor 2 | missing of Auto stop of RO& EDI trains in case of tank high level | auto stop implementation in WTP DSC logic | Moderate | Jul 30, 2020 | include in WTP improvemenet plan and implementaion ongoing by ROCKWELL AUTOMATION | Closed | 06GCL11CL001/06GCK21CL001/06GCK31CL001
47 | Noor 2 | deviations of RO(RO1&RO2) flow metres (feed ,permeat ,reject)according to reference third part . | Calibration /change of folw meters . | High | Jul 30, 2020 | PO for procurement of new flow meters had been send by I&C team | Closed
48 | Noor 2 | missing of temperature transmitor in ilnet RO system | install online temperature sensor | High | Jul 30, 2020 | defect was send to EPC and procuremen of temperature sensor | Closed
49 | Noor 2 | missing of Auto sotp of RO train in case of hight conductivity and ORP | auto stop implementation in WTP DSC logic | Moderate | Jul 30, 2020 | include in WTP improvemenet plan and implementaion ongoing by ROCKWELL AUTOMATION | Closed | 06GCK40CQ002/06GCK40CQ002
50 | Noor 2 | missing of auto startup of clear water pump and waste water pump in case of pond hight level | auto startup implementation in WTP DSC logic | Moderate | Jul 30, 2020 | include in WTP improvemenet plan and implementaion ongoing by ROCKWELL AUTOMATION | Closed | 06UGP12CL001/06GMC11CL001
51 | Noor 2 | Recercylating of MMF&ACF (Turb>30NTU,PH=8.2 also flash water charged with entrained sand and charbon in backflushing pond after one month of operation ) backfushing water to inlet of pretreatment | EPC should review the water balance and provide the specification of flash water to recyrculate to inlet pretreatment . | Moderate | Jul 30, 2020 | installation of sludge pump in MMF backfush pond and recycling water to waste water pond | Closed
52 | Noor 2 | manual Amonia dosing in feed water and mackup condensate water | implementation of Amonia Auto dosing | High | Jul 30, 2020 | implementation of Auto dosing logic by SCHNEIDER ongoig after provid of logic by EPC | Closed | 25QCB13/14/15/16/AP001
53 | Noor 2 | oil analyzers require calibration and manuals including all standard parts | EPC should calibrate and provide all required records | High | Jul 30, 2020 | third part offre was send to APO and still waiting to provid budget for carrig out the calibration and commisioning of oil analyzers | Open
54 | Noor 2 | lack of laboratory quality manual according to ISO 17025 guidelines | training the staff and implementation of QMS ISO 17025 | Moderate | Jul 30, 2020 | work instruction of water analyzers was developed and manual in progress | Closed | REP-2019-NRR002
55 | Noor 2 | Lack of HVAC water chemistry manual include (water quality limit and corrosion inhibitor dosing system)with the procedure of replacing water in case of exedences | EPC should provid HVAC water chemistry manual | Moderate | Jul 30, 2020 | EPC communicate insufficient document ,development of guidelines refering to international references | Closed
56 | Noor 2 | lack of sludge treatment procedure (floculant dosing,startup conditions …)was communicated by EPC | EPC should provide sludge treatment procedure | Moderate | Jul 30, 2020 | defect of missing records was send to EPC and development of WI refereing to international reference | Closed | REP-2019-NRR002
57 | Noor 2 | lack of RO1 and RO2 simulation datasheet | EPC should provide simulation datasheet of RO system | Moderate | Jul 30, 2020 | simulation of RO performance by HYDRANAUTIC company | Closed | REP-2019-NRR002
58 | Noor 2 | SWAS circuit damaged with HTF (folw meters , cartidge filter &boxes, plastic piplienes ...) | procureement of SWAS spart parts | High | Jul 30, 2020 | installation of new cartidge filtre and fow mettres | Closed
59 | Noor 2 | lack of Auxialary boiler chemistry manual include (feed water ,drum ,superheated steam ) | EPC should provide Auxialary boiler chemistry manual | High | Jul 30, 2020 | defect was send to EPC for providing manual and development of WI based on international references | Closed
60 | Noor 2 | lack of WSAC chemistry manual include (feed water ,bassin guidelines ) | EPC should provide WSAC chemistry manual include (feed water ,bassin guidelines ) | High | Jul 30, 2020 | EPC communicate insufficient document ,development of guidelines refering to international references | Closed
61 | Noor 2 | lack of Corrosion inhibitor dosing system for CCCW | install of Corrosion inhibitor dosing system for CCCW. | Moderate | Jul 30, 2020 | EPC provid manaual injection system | Closed | .
62 | Noor 2 | lack Alkalin &Acid dosing system for effluent neutralization in monitoring pit | install of Alkalin &Acid dosing system for effluent neutralization in monitoring pit | Moderate | Jul 30, 2020 | procurement of dosing pump for injection | Closed
63 | Noor 2 | missing of Regulation loop bettween Bassin conductivity and blowdouwn regulation valve | EPC should implement the regulation loop bettween Bassin conductivity and blowdouwn regulation valve | Moderate | Jul 30, 2020 | implementation of Auto dosing logic by SCHNEIDER ongoig after provid of logic by EPC | Closed | 25PGD10AC001
64 | Noor 2 | missing of Regulation loop bettween corrosion inhibitor pump and mackup valve | EPC should implemnt the regulation loop | High | Jul 30, 2020 | defect was send to EPC and waiting to provid and implement logic | Open | 25PGD10AC001
65 | Noor 2 | missing of Regulation loop bettween sulfiric acid pump and bassin pH | EPC should implemnt the regulation loop | Moderate | Jul 30, 2020 | defect was send to EPC and waiting to provid and implement logic | Open | 25PGD10AC001
66 | Noor 2 | lack of SWC preservation in case of long shtdonwn | EPC should provide preservation procedure | Moderate | Jul 30, 2020 | development of our own work instruction based on international requirement | Closed
67 | Noor 2 | lack of antifreez in CCCW ,risk of water freezing in negative temperatures | EPC should provid procedure and product specification | Low | Jul 30, 2020 | accrording to third part and EPC antifrezz based on city weather not recommended | Closed
68 | Noor 2 | soduim analyzer of condensate water require spart part ( sensor …..) | Change sodium reference sensor | High | Jan 3, 2020 | soduim reference sensor was changed and analyzer is operational | Closed
69 | Noor 2 | SGS sencond line monitoring with online analyzers | Install a new monitoring online for SGS B | High | Aug 30, 2020 | SWAN offer for installation of second line was shared with APO and EPC and still waiting budget approval for implementation | Open | 25QUC/25QUB
70 | Noor 2 | pH ,conductivity Online analyzer not operationel | Change the pH and conductivity sensors damaged | High | Jul 30, 2020 | spart part received and IC team will install new PH and conductivity sensor | Closed | 25GHC50CA001/002/003
71 | Noor 2 | pH ,conductivity Online instrument of WSAC not operationel | Change the pH and conductivity sensors damaged | High | Jul 30, 2020 | spart part received and IC team will install new PH and conductivity sensor | Closed | 25GBK20CA001/002 - 25PGD10CA001
72 | Noor 2 | Visual inspection ACC, SGS heat exchangers | ADQUIM should provid the Visual inspection report | High | Apr 15, 2020 | visual inspection was done and report received | Closed | NOOR2 power block
73 | Noor 2 | HTF leaks in evaporator B | EPC should open and repair evaporator tubes leaks | High | Feb 20, 2020 | evaporator leakage was repaired during forced outage | Closed | 25HAD20AC001
74 | PetroRabigh-RAWEC | The ammonia dosing control is set in auto mode.The dosing target is 7 - 8 µS/cm (taken from Chemical Quality Plan NRPC/OPK/SP-001). Note that the steam-off-taker also dictates the target: maximum pH 9.6 | Assess increasing the ammonia dosing control target to 9 – 10 µS/cm. This range corresponds to pH 9.52 – 9.57. | Low | Apr 29, 2023 |  | Open
75 | PetroRabigh-RAWEC | With a relatively stable operating regime, the phosphate concentration decreases due to water loss from the drum/waterwall system. Through blowdown and/or through passing valves | 01. Address the passing valves | Moderate | Apr 30, 2023 |  | Open
76 | PetroRabigh-RAWEC | the ammonia concen-tration remaining in the drums increases from 430 ppb to 630 ppb. The resulting drum water pH without adding TSP will already be ± 9.25. Hence only a minor TSP dosage is needed to lift the pH over the 9.30 threshold | Modify the TSP dosing. Don't allow the pH to drop too low, and prevent the pH from increasing too high. This can be done by lowering the TSP concentration and/or modifying the dosing pump flow rate | Moderate | Apr 30, 2023 |  | Open
77 | PetroRabigh-RAWEC | Online analyzers | Install HHP steam header online CC measurements | High | Apr 30, 2023 | P | Open
78 | PetroRabigh-RAWEC | Online analyzers | Install HHP steam header online sodium measurements | High | Apr 30, 2023 |  | Open
79 | PetroRabigh-RAWEC | The silica targets for HHP FW (< 1000 ppb – target < 300 ppb), DEA FW (< 300 ppb – target < 5 ppb) are inconsistent and incorrect | Modify the water and steam purity specifications. Use the table in Annex 1 as a reference. | Moderate | May 10, 2023 |  | Open
80 | PetroRabigh-RAWEC | International organizations and turbine suppliers include online steam sodium monitoring as a key measurement. Moreover, most turbine suppliers have steam CC and sodium as "hold" parameters for admit-ting steam to the turbine. Steam can only be admitted when CC and sodium reach a specific target. | Add steam purity specifications for steam turbine startup. | Moderate | May 10, 2023 |  | Open
81 | PetroRabigh-RAWEC | The installed CC analyzers are programmed with an incorrect tempera-ture compensation factor. | Program a 1.4 % slope temperature compensation for the CC measurements. | Low | Apr 15, 2023 |  | Open
82 | PetroRabigh-RAWEC | SWAS - Online PH and SC calibration procedure | Abandon the lab pH and SC measurements on grab sam-ples using the lab instrument and implement a validation procedure as described/shared by consultant | High | May 20, 2023 |  | Open
83 | PetroRabigh-RAWEC | Replacement of PH electrodes | Assess the lifetime of the pH electrodes and replace them when needed. | Moderate | Apr 30, 2023 |  | Open
84 | QIPP | Calibration – Validation Program The I&C department does calibrations according to a fixed schedule or on the demand of the chemistry department | Abandon the lab pH and SC measurements on grab samples using the lab instrument and implement a validation procedure as de-scribed | High | Apr 10, 2023 | instrument's procurment is in progress | Open
85 | QIPP | Calibratiob Analyzers & Measuring pH and SC using the lab equipment has low added value | Upgrade the instrument calibration procedures. | High | May 30, 2023 |  | Open
86 | QIPP | PH and Sodium electrode replacement | Assess the lifetime of the pH and sodium electrodes and replace them when needed | High | Apr 30, 2023 | requested additional budget was submitted by IMD head | Open
87 | QIPP | SC and CC online analyzers calibration | Do not perform a calibration on the SC and CC analyzers using the available standard solution | High | Apr 10, 2023 |  | Closed
88 | QIPP | Grab samples are taken daily for pH and SC. Silica, phosphate, and iron are measured 3 times per week. The results are entered in the "HRSG & RO Lab" spreadsheet. | Implement an Instrument Standard Operating Procedure | High | Apr 30, 2023 | instrument's procurment is in progress | Open
89 | QIPP | Document "NQIPP-OPA-SP-005 QIPP Condenser Leak Handling Proce-dure" provides the action plan in case of seawater ingress | Further improve document "NQIPP-OPA-SP-005 QIPP Con-denser Leak Handling Procedure" by (1) making the relationship be-tween condensate sodium (once installed) and condensate CC essential in establishing the presence of seawater ingress | Moderate | May 31, 2023 |  | Open
90 | R2IPP | The ammonia dosing control is set at manual. Feedwater SC and pH are monitored and manual adjustments to the ammonia pump stroke are made. | Set the ammonia dosing control to automatic (fixed pump stroke and - frequency control for adjustment Use 10(20-30)QUA01CQ003 as ammonia dosing control parameter – set target SC at 11 µS/cm | High | Apr 30, 2023 | MOC created on 02.02.2023 (N#280322) and waiting for I&C | Open
91 | R2IPP | TSP is dosed in the LP, IP, and HP drums. | Document NR2IPP/OPK/WI-025 Rev.00 Plant Chemistry Management, paragraph 9.4.2. Boiler Drum Phosphate Dosing, only mentions dosing of TSP in the LP and HP drum. This is not consistent with on-site find-ings. Dosing TSP in the HP drum is, during trouble-free operation, not re-quired when the correct feedwater pH is attained | Low | May 30, 2023 | Daily HRSG steam & water analysis report updated /awaiting for management approvel | Closed
92 | R2IPP | Manual dosing control based on the pH and on the phosphate concentra-tion measured in a grab sample. | The pH ranges are set too high. During trouble-free operation, an LP drum water pH of 9.4 is sufficient; a pH of 9.3 for the IP drum.1 ppm PO4 in the IP drum should bring the pH to 9.3. Adjust the TSP dosing in the LP and IP drums. | Moderate | May 30, 2023 | Daily HRSG steam & water analysis report updated /awaiting for management approvel | Closed
93 | R2IPP | The Daily Station Chemistry Report gives cation conductivity (CC), so-dium, and silica limits.No condensate sodium target is given, nor is condensate (QUCCQ01) sodium measured ( | The limits used are not fully consistent with international best prac-tices: the CC target is set too low (unnecessarily low), the sodium limit is not adapted to recent guidance values (Dooley et al., 2006; Petrova, 2013). Modify the guidance values according to the table in Annex 1. Priority: Moderate. | Moderate | May 30, 2023 | Daily HRSG steam & water analysis report updated /awaiting for management approvel /Group #1 and Group #3 ,Lp steam sodium analyzer exchanged to CEP water | Closed
94 | R2IPP | For power plants cooled by saline waters, particularly if there is no pro-vision for condensate polishing, condensate sodium measurements must be included in the Key Instrumentation list | Purchase one new sodium analyzer for each block asap and install it on the condensate sample. On the sodium analyzer (QUH09CQ003), measuring LP SATST and LP SHST sodium, replace the LP SATST with the unit's condensate sam-ple. | High | Apr 30, 2023 | MOC created on 02.02.2023 (N#280337) and waiting for I&C | Open
95 | R2IPP | Drum water silica limits are pressure dependent. The maximum allowed silica concentration in the drum water decreases with increasing pres-sure. Operating outside the threshold increases the risk of too high sil-ica in the steam. The drum silica limits are set too high (5 ppm). | Modify the drum silica limits according to the table in | Moderate | May 30, 2023 | updated/awaiting for mangment approval | Closed
96 | R2IPP | Measuring the Dissolved Oxygen (DO) (noted as QUH07CQ002 and mentioned on the Daily Station Chemistry Report) in drum waters has no added value. | Consider drum water DO measurements as obsolete | Low | Apr 20, 2023 | Updating procedure ,updating daily water chemistry report | Open
97 | R2IPP | For power plants cooled by saline waters, particularly if there is no pro-vision for condensate polishing, condensate sodium measurements must be included in the Key Instrumentation listI | Install sodium analyzers on the condensate | High | Apr 30, 2023 | duplicated | Closed
98 | R2IPP | pH analyzers: The pH analyzers do not all have the Process Tempera-ture Compensation (PTC) activated. pH measurements without the PTC activated do not give the correct result when the sample temperature deviates from 25 °C. | Adjust the PTC setting on the pH analyzers to -0.029. See transmitter manual chapter | High | Apr 30, 2023 | related with I &C | Open
99 | R2IPP | Conductivity analyzers: For some analyzers, the temperature compen-sation is not adapted to the nature of the sample | Verify and correct, when needed the temperature compen-sation on all SC (set to slope 2%) and CC (set to "cation") transmitters. See transmitter manual chapters 6.4.8 and .9 | High | Apr 30, 2023 | related with I &C | Open
100 | R2IPP | Calibration – Validation Program | Abandon the lab pH and SC measurements on grab sam-ples using the lab instrument and implement a validation procedure as described in Annex 3. (and in our document3 MEM-2022-NHQ231137). | High | Apr 30, 2023 | PR 3200025363 raised for potable SWAN analyser | Closed
101 | R2IPP | Calibration – Validation Program | Do not perform a calibration on the SC and CC analyzers using the available standard solution. Instead, when validation exe-cuted according to the procedure | High | Apr 30, 2023 | PR 3200025363 raised for potable SWAN analyser | Closed
102 | R2IPP | Laboratory work is currently focused on sampling and analyzing. This methodology has a low added value as no documented action plan is related to the outcome of the analysis | Implement a validation procedure as given in Annex 3 and upgrade the reporting. Link operational procedures to detect upset con-ditions. | High | Apr 30, 2023 | PR to be raise under PC cost | Open
103 | R2IPP | No validation or online instrument QA/QC program is available. | Implement an Instrument Standard Operating Procedure | High | Apr 30, 2023 | I & C will be update the online analyzer validation/calibration | Open
104 | R2IPP | When evaluating the iron content, it must be understood that the filter AT001 can retain iron oxide particles; thus, the total iron content can-not be correctly measured. | Use the available Inductively Coupled Plasma analyzer for the iron measurements. | Low | Apr 25, 2023 | ICAP instrument need to service & PR 3200014950 raised and still with SC under process | Closed
105 | R2IPP | No formal procedures or work instructions based on well-defined target and/or alarm levels are yet available | Develop chemistry-related operational procedures based on the targets and AL's given in Annex 1. These procedures should cover, among others, a condenser leak action plan, blowdown opera-tion, drum pressure modifications, manual interventions on TSP dosing system, verification of the automatically controlled ammonia dosing system, and instrument calibration and maintenance. | Moderate | May 10, 2023 | need to create new SOP | Open
106 | SALALAH | • The CPD sample is the correct sample to control the ammonia dosing.The controlling parameter is CPD pH (10QUC10AK006). The dosing control was in manual mode. • Target pH range: 9.2 – 9.8; target 9.4. | R1.3.1. Ammonia dosing to the condensate: Set the ammonia dosing control to automatic (pump stroke at 50% - frequency controlled dosing). R1.3.2. Change to dosing control parameter to 10(20)QUC10AK005 (CPD after chemical dosing SC) R1.3.3. Set target setpoint to 17 µS/cm (range 14-26 µS/cm) | High | Nov 1, 2019 |  | Closed | REP-2019-NSL001
107 | SALALAH | The used dosing target range of 0.3 – 7 ppm was set by the contractor. | R2.1.1. Limit the TSP dosing to a maximum of 3 ppm PO4. | High | Nov 1, 2019 |  | Closed | REP-2019-NSL001
108 | SALALAH | TSP is dosed in the HP drum to a maximum of ± 3ppm PO4. | R2.2.1. Stop dosing of TSP in the HP drums. | High | Nov 1, 2019 |  | Closed | REP-2019-NSL001
109 | SALALAH | Temperature compensation | Process Temperature compensation must be activated on all pH analyzers | Moderate | Dec 31, 2019 | In progress - Need to dsicuss with the OEM to implement the TC factor | Closed | REP-2019-NSL001
110 | SALALAH | Alaram and Alert | Configure the site Alarm and Alert levels as per the report | Moderate | Dec 31, 2019 | In progress | Closed | REP-2019-NSL001
111 | SALALAH | No written procedures available addressing upset conditions | Write action program for high/low pH condensate, high/low pH IP drum water, high silica, high CC in HP drum water | Moderate | Dec 31, 2019 | In progress | Closed | REP-2019-NSL001
112 | SALALAH | No written procedures available addressing instrument nonconformity | Write procedure objectively outlining when an instrument needs maintenance or calibration by I&C department | Moderate | Dec 31, 2019 | In progress | Closed | REP-2019-NSL001
113 | SALALAH | Iron analysis | Set up a monthly sampling program for iron (together with other NOMAC Oman plants). Samples must be sent to an external laboratory for iron analysis using either Atomic Absorption Spectrometry or Inductively Coupled Plasma. This allows for the build-up of a correct historical data-base to be used in case of tube failures. | Low | Dec 31, 2019 | In progress | Closed | REP-2019-NSL001
114 | Shuqaiq | Eliminox (Carbohydrazide solution) is dosed in the condensate line down-stream of the CPD sample | Stop the oxygen scavenger dosing | High | Apr 20, 2023 | 1)Need guideline to phase out dosing of Oxygen scavenger 2)MOC with PC approval will be initiated as per guideline | Open
115 | Shuqaiq | condensate sodium measurements must be included in the Key Instrumentation list. | Procure and in-stall online condensate sodium analyzers | High | May 30, 2023 | 1)Need guideline with suggested OEM/Vendor, installation location, range with alarm limits etc. 2)MOC with PC approval will be initiated as per guideline 3)Procurement action will be taken after MOC approval 4)Discussion with MHPS for intigration of new analysers signals with DCS | Open
116 | Shuqaiq | On all 3 sampling racks, sample flow problems were seen. These problems compromise the reliability of the measurement, notably on the cation conductivity measurements. | Set up a critical inspection program for the sampling system—Refurbish parts when re-quired. | Low | May 30, 2023 | Aknowledged | Closed
117 | Shuqaiq | spare resin columns for cation conductivity are unavailable, and the mainte-nance/calibration program needs to be up-dated. | Implement an Instrument Standard Operating Procedure that includes a validation program using portable reference instruments. Correct the temperature compensation settings on the pH and conductivity analyzers where needed | Moderate | May 20, 2023 | Spare Resin columns for cation conductivity to the tune of 50% has been precured and are available with us. Only need to Correct the temperature compensation settings on the pH and conductivity analysers wherever needed. | Open
118 | Shuqaiq | Operational knowledge of the water steam cy-cle chemistry is present. | Continue consolidating the knowledge in a chemistry manual and related work instruc-tions. | Moderate | May 20, 2023 | Aknowledged | Closed
119 | Shuqaiq | Dosing control - Continue consolidating the knowledge in a chemistry manual and related work instruc-tions. | Assess using the deaerator inlet SC as the ammonia dosing control parameter. | Low | Jun 30, 2023 | Our ammonia dosing pumps are following the dearator inlet/Outlet Conductivity as per Operating condition, require further recommendation | Open
120 | Shuqaiq | A DSP – TSP mixture is dosed in the drums | The product is not OK.Replace the DSP – TSP mixture with TSP only Dosing phosphate in the HP drum is, during trouble-free operation, not required when the correct feedwater pH is attained | Moderate | May 30, 2023 | Aknowledged | Closed
121 | Shuqaiq | Manual dosing control based on the pH and on phosphate concentration measured in a grab sample.-Dosing control | If TSP is used, limit the concentration to 1 ppm. | Moderate | May 30, 2023 | Aknowledged | Closed
122 | Shuqaiq | Related to water and steam purity, the Power Water Analysis table from the document "SQIWPP Plan Product Quality Assurance & Control (Power) NSQIWPP/QAC/PL-002" gives cation conductivity (CC), silica, and Dissolved Oxygen (DO) limits. Also, iron, copper, and chloride lim-its are given. | The limits used are not entirely consistent with international best prac-tices: Assess implementing the purity limits given in the reference table in Annex 1. Technical Guidance Values. Assess using Alarm Levels | Moderate | Jun 10, 2023 | Logic modification by DCS OEM with approval from PC | Open
123 | Shuqaiq | CC and silica are measured at the correct locations Sodium in the condensate and Superheated Steam is not measured. | Online condensate sodium measurement is a must. | High | Apr 30, 2023 | 1)Need guideline with suggested OEM/Vendor, installation location, range with alarm limits etc. 2)MOC with PC approval will be initiated as per guideline 3)Procurement action will be taken after MOC approval 4)Discussion with MHPS for intigration of new analysers signals with DCS | Open
124 | Shuqaiq | The pH analyzers have the Process Temperature Com-pensation (PTC) activated | Adjust the PTC setting on the pH analyzers to -0.33/ 10°C Attention: check the manual as some Yokogawa EXA 402 series use a compensation factor per °C. In this case, -0.033 needs to be pro-grammed. | High | Apr 30, 2023 | Under Progress | Open
125 | Shuqaiq | The I&C department does calibrations according to a fixed schedule or on the demand of the chemistry department. No pH, DO, or silica calibration records were available. | Abandon the lab pH and SC measurements on grab sam-ples using the lab instrument and implement a validation procedure. Implement the Instrument Standard Operating Procedure | High | Apr 30, 2023 | To be followed | Closed
126 | Shuqaiq | Document "NSQIWPP-QAC-PL-002_r3 Product Quality Assurance & Control (Power)" table "6.2.4 Boiler Water/ Steam" lists limits and ac-tions to be taken in case of deviation | Develop chemistry-related operational procedures based on the targets and AL's given in Annex 1. These procedures should cover, among others, a condenser leak action plan, blowdown opera-tion, drum pressure modifications, verification of the automatically con-trolled ammonia dosing system, and instrument calibration and mainte-nance. | Moderate | May 30, 2023 | Meetings will be scheduled to share modified procedures for awareness of concerned employees. | Closed
127 | Shuqaiq | Documents "P0-MN-AX-HAC-WE1-00001_r1_Boiler APS Startup" and "P0-MN-AX-MAA-WE1-00001_r1_Turbine APS Startup" not refer to chemistry. | Consult with the turbine supplier (and others) on the spec-ifications for steam admission. | Low | May 10, 2023 | To Be Consulted with MHI | Open
128 | Sohar 3 | Ammonia dosing is best on auto control given the shifting load | Ammonia dosing pump fixed stroke - Frequency controlled | High | Nov 11, 2019 |  | Closed | REP-2019-NSH002
129 | Sohar 3 | Dosing TSP | No TSP in the HP drums. | High | Nov 11, 2019 |  | Closed | REP-2019-NSH002
130 | Sohar 3 | Process Temperature compensation must be activated on all pH analyzers | Modify instrument settings | High | Nov 11, 2019 |  | Closed | REP-2019-NSH002
131 | Sohar 3 | Some instruments analog output signals need modifications | Modify instrument analog output signal according report REP-2019-XXX002 | High | Nov 11, 2019 |  | Closed | REP-2019-NSH002
132 | Sohar 3 | Reference Instruments (SWAN) | Purchase reference instruments for conductivity and pH | Low | Dec 31, 2019 | Already i Raised the PR(0010032446) | Closed | REP-2019-NSH002
133 | Sohar 3 | Upgrade the lab work | Use of reference instruments for conductivity and pH instead of grab sample analysis | Moderate | Nov 30, 2019 | Implemented | Closed | REP-2019-NSH002
134 | Sohar 3 | Instrument validation program | Implement the Instrument Standard Operating Procedures | Moderate | Dec 30, 2019 | In-Progress | Open | REP-2019-NSH002
135 | Sohar 3 | NEVER LET ANY CC RESIN COLUMN BE EXHAUSTED!!! | Write a rigorous procedure to avoid operating with exhausted resin columns | Moderate | Nov 20, 2019 |  | Closed | REP-2019-NSH002
136 | Sohar 3 | No written procedures available addressing upset conditions | Write action program for high/low pH condensate, high/low pH LP/IP drum water, high silica, high CC in HP drum water | Moderate | Jan 10, 2020 | Submitted for Approval | Closed | REP-2019-NSH002
137 | Sohar 3 | No written procedures available addressing instrument nonconformity | Write procedure objectively outlining when an instrument needs maintenance or calibration by I&C department | Moderate | Jan 10, 2020 | Submitted for Approval | Closed | REP-2019-NSH002